Title

A Comprehensive Testing Protocol for Macro-scale Mechanical Characterization of Knee Articular Cartilage with Documented Experimental Repeatability

Authors

Snehal Chokhandre1
Ahmet Erdemir1

Affiliations

1. Department of Biomedical Engineering, Lerner Research Institute, Cleveland Clinic, Cleveland, OH.

Corresponding author

Ahmet Erdemir, PhD
9500 Euclid Ave, ND-20,
Cleveland, OH, 44195, USA.
 erdemira@ccf.org

Abstract

Articular cartilage mechanics has been extensively studied with various approaches and mechanical characterization strategies. However testing protocols can be highly varying and difficult to reproduce, particularly for specimen-specific analyses. Detailed knowledge of testing protocols is important for reliable use in concordant finite element analyses. This study presents a detailed, robust procedure for cartilage testing—with multiple regions and per specimen repeatability data. Samples were taken from femur, tibia and patella of a human cadaver knee and tested in unconfined compression, confined compression and uniaxial tension. Each test was repeated three times. The testing protocols provide elastic and time dependent characterization data. Results, for example equilibrium modulus of 0.28 (0.0024) MPa for patella under unconfined compression indicate that variability is well controlled and that protocol(s) presented here can generate repeatable specimen-specific data. As per the authors' knowledge this is the first study to report in-depth uncertainty assessment of the experimental procedures for multi-region knee cartilage characterization.

Graphical Abstract

See graphical abstract document

Highlights

- A well-defined testing protocol was developed for mechanical characterization of articular cartilage.
- Major test types for typical characterization were included: confined and unconfined compression and, uniaxial tension including multi-step stress relaxation.
- Repeatability of the testing procedures and resultant tissue behavior was assessed by performing each test on the same sample three times.
- The protocols developed for material characterization produced repeatable data with location dependent variability and variability assessments within the same sample-test combination.
- Study is first of its kind to document uncertainties of experimental procedures for multi-region cartilage characterization.
- Documented repeatability of experimental procedures can improve delineation of protocol uncertainties from regional versus biological variations.

Abbreviations
AIM: Average instantaneous modulus, AEM, Average equilibrium modulus, UC: Unconfined compression, CC: Confined compression, UT: Uniaxial tension.

Keywords
Human knee, articular cartilage, experimental repeatability.

1. Introduction

Finite element (FE) analysis has emerged as a critical tool for understanding the mechanics of the overall knee joint and its sub-structures, including articular cartilage1–5. Variations in linear and nonlinear macro and multi-scale cartilage material properties as well as specimen-specific and regional variations can have significant impact on FE model analyses and predictions 6–8. Determination of these properties is challenging due to the cartilage’s anisotropic, non-homogeneous, non-linear behavior as well as multi-phasic and time dependent properties. Most FE studies rely on material parameter information from literature7,9,10. Reconciling differences in literature derived properties reported by various experimental studies11–15 and selecting suitable values for a new FE model can thus be difficult. Often the protocols for testing are summarized and can be very different between studies. Furthermore, reported values may not represent the specimen-specific properties for which a new individualized FE model is being developed. These discrepancies may affect the model calculations and predictions6,8,16. It is important to note that even if exact specimen-specific properties are not available for FE model development, availability of studies that can provide relevant information to make informed decisions are of paramount importance, e.g. ranges of possible material properties for a given demographic, location dependent relationship between material properties, etc. Apart from the applications to FE modeling, detailed descriptions and standardization of characterization protocols encompassing a wide range of potential mechanical behaviors can provide insights into the tissue mechanics as a whole.

Three primary methods currently used for characterizing cartilage material properties are confined compression, unconfined compression, and indentation15,17–25. Tensile properties of the superficial layer of the cartilage, whole depth and to a lesser extent regional properties and depth variations have also been obtained26–31. There is large variability in experimental protocols used for characterizing these nonlinear elastic and time dependent properties (see Appendix A), in addition to a wide range donor characteristics, e.g., age, overall health/activity, etc., and species of interest.

The types of tests performed often depends on the needs of the particular study. For determining multiple material model parameters, multiple types of tissue tests may be needed. A study by Ahsanizadeh and Li32 indicated that in addition to confined compression, both simple tensile and multi-step stress relaxation tests are required to obtain sufficient material parameters for an anisotropic fibril reinforced material model. Combining of unconfined and confined compression tests allows assessment of effects of unidirectional deformations (viz. confined compression) and unidirectional loading (viz. unconfined compression) providing basic properties such as Young’s modulus, aggregate modulus, hydrostatic permeability, and Poisson’s ratio. Multi-step stress-relaxation tests can provide data for non-linear relaxation models. Preconditioning methods (whether performed or not), preconditioning percent strain, and number of cycles also affect the tissue behavior in subsequent stages of the test33–35. Additionally, the target strains and strain rates used in non-linear elastic and time dependent characterization need to be within non-destructive ranges.

Many pragmatic and practical testing procedures need to be considered such as: required target strains, target strain rates, location of samples, dimensions and orientation of samples, number of freeze thaw cycles, test environment (immersing fluid, temperature of environment), number of preconditioning cycles, preconditioning strains and rates, sample preparation, potential sample degradation, clamping effects and orientation, stability and permeability of the filter in confined compression, capabilities and performance of the testing equipment, sample dimension measurement errors, and so on can all effect testing outcomes. Acquisition of detailed material properties, whether as a quality metric for tissue health or to use in computational models, requires standardized and repeatable documentation and evaluation of experimental protocols. Understanding uncertainties associated with testing protocol is thus imperative when attempting to separate actual regional and biological variations in mechanical properties of cartilage from potential specimen preparation and measurement errors.

The two main goals of this study were: i) to present a comprehensive testing strategy for macro-scale material characterization of articular cartilage with explicitly documented testing procedures, and ii) to quantify repeatability as an indicator of testing reproducibility and experimentation errors. The detailed methods/protocols and data provided here are anticipated to be useful for experimental researchers developing new testing strategies and ultimately for knowledge-based comprehensive FE models of the whole knee. Pragmatics of testing procedures and protocol development considerations are also described.

2. Materials and Methods

2.1 Specimen overview
The study was conducted on a cadaver left knee from a 25 years old female donor (caucasian, body mass index: 22.8, height: 1.73 m, weight: 68 kg). The specimen, labeled oks003, is part of the Open Knee(s) cohort, an initiative to develop specimen-specific finite element models of knees (https://simtk.org/projects/openknee). A radiological assessment on X-ray image provided by the vendor (Anatomy Gifts Registry, Hanover, MD), was performed by a surgeon to rule out any obvious injury or disease before the specimen was approved for procurement. Serological testing documentation was also provided by the vendor before approval.

2.2 Sample location and preparation
Cartilage strips were carefully separated from bones using a scalpel from the load bearing regions of the knee. Care was taken to make sure the strips were large enough to accommodate at least one compression and one tensile sample. It is important to note that the cartilage samples had been frozen and thawed twice before testing as part of Open Knee(s) protocols. First thawing cycle was for joint imaging and mechanical testing (not discussed, out of scope for this study) after procured frozen from vendor and subsequent dissection for separating cartilage strips from the bones after which the strips were frozen (-20 °C). The second time, each strip was thawed once and both compression and tension samples were prepared and either tested right away or kept in the fridge for later testing.

Rectangular cartilage strips were taken from the lateral and medial femoral condyles, lateral and medial tibial plateaus, trochlear groove and patella (Figure 1). All cartilage strips were allowed to stabilize in phosphate buffered saline (PBS, 10x solution, diluted with de-ionized water) before the samples were punched out. One full thickness compression sample was punched out from each strip using a custom punch of 5 mm in diameter. Initially, a full thickness planar dumbbell shaped tensile sample (5 mm by 1 mm test area) was also punched out from each of the six strips. The longer dimension was in the direction of fibers in the superficial layer36. Once this sample was punched out, a thin uniform thickness sample closer to the superficial layer was prepared using a vibratome (Leica VT1000s, Leica Camera AG, Germany). Uniform thickness compression samples were also prepared using the vibratome. The punched cylindrical samples were fixed using minimal tissue adhesive on the sample plate on the vibratome and using a vibrating carbon steel blade (moving at 1 mm/s, 2 mm amplitude), minimal tissue from the top surface was carefully cut to maintain as much thickness as possible. For tensile samples, since we were interested in obtaining the tissue close to the superficial layer, care was taken to cut the bottom of the full thickness dumbbell sample manually to have a flat sample first. This sample was then placed on the vibratome sample plate using tissue adhesive and the top layer was carefully excised to be used for tensile testing. Goal was to obtain the tissue as close to the superficial region as possible.

Once the samples were ready, they were kept immersed in PBS until further procedures were conducted. This typically did not exceed 1 hour. In some cases, compression sample was punched out when the strip was still frozen and the strip was stored back in the freezer until a tensile sample was punched out later. A total of 12 samples were prepared and tested over the course of the study (6 compression samples and 6 tensile samples).

2.3 Sample dimension measurement
A custom optical thickness measurement system (OTMS) was built to facilitate non-contact measurement of sample thicknesses (Figure 2). A camera (FlyCapture, FMVU-13S2C-CS, FLIR Systems Inc, Wilsonville, OR) was used to manually take pictures of the samples placed on the OTMS platform. A ruler was placed next to the sample to calibrate pixel measurements to physical units (mm). A Python script was developed to manually pick multiple thickness measurements on the image of the sample, calculate average and standard deviation and save the original image, image with measurements, and an XML file with all the measurements. For this study, 10 thickness measurements were taken for each of the sample. Ten measurements manually taken on the sample image ensured that as much as of the sample thickness is covered in the measurements for both the compression and tensile samples. As thickness was measured each time before the sample was tested under compression test, any error associated with repeated measurements would reflect any changes in geometry between tests and also errors due to manual measurements. Thickness of tensile samples was measured once before the test set. Width for tensile samples was assumed 1 mm and diameter for compression samples was assumed 5 mm for all calculations based on punch dimensions. Tensile sample reference length was measured by application of 10g load on the tensile sample during tests (see details in following sections).

2.4 Testing equipment and testing environment
All the tests were conducted on a Mach-1 mechanical testing machine (Biomomentum Inc, Laval, Québec, Canada), with a 10 kg load cell (resolution: 0.5 g). Three types of tests were conducted: confined compression (CC), unconfined compression (UC) and uniaxial tension (T) (Figure 3). Each of the 6 compression samples was tested under both confined and unconfined compression. Custom test chambers were designed for each of the test type. For unconfined compression, the sample simply sat on an acrylic plate that attached to the bottom of the test chamber and a flat indenter (with a diameter larger than the sample) was used to compress the sample. For confined compression, a compression chamber made of two parallel acrylic plates with a cylindrical cavity on the top plate was designed. Once the sample was placed in the compression chamber, a stainless steel filter (5 mm diameter, 5 μm pore size, thickness 1 mm) was placed on top of the sample. A 3 mm diameter indenter rod was then used to compress the sample by pushing on the filter. Custom sawtooth stainless steel clamps were used to grip the tensile samples. Additionally, tissue adhesive was used to ensure that the samples would not slip out of the grips. All the tests were conducted at room temperature and the samples were immersed in PBS for the entire duration of the test. The tensile samples were marked with a marker pen and a ruler was placed in the bath parallel to the tensile sample for potential optical strain measurement once the images were captured using the camera system of Mach-1.

2.5 Testing protocol
A detailed description of the protocol is provided in Table 1 with all the relevant parameters and justifications for using those parameters. Figure 4 provides an illustration of the protocol. First, a ramp load was applied to maximum target strain of 15% to assess the tissue properties before any preconditioning was done. This ensured the capture of potential effects due to pre-testing steps (sample acquisition, preparation etc). Both the pre- and post-preconditioning ramp loads allowed the estimation of elastic properties before and after preconditioning. Establishing a known and repeatable reference state for subsequent reliable characterization is important. A total of 1000 cycles was determined from pilot tests that demonstrated sufficient cycles for repeatable behavior. A study done by Cheng et al33 indicated that preconditioning should be conducted to a maximum target strain at which the later tests are going to be conducted. Incidentally, using the maximum target strains for preconditioning also helped improve repeatability in the following stages of the given test. As the entire test was repeated including moving the sample out of the test environment to a refrigerator for two hours, and measuring thickness etc again, the initial state of the tissue was assumed to be reset and preconditioning established the initial state again at each test in a given set. The stress relaxation data provided means to study time dependent properties of the tissue. Compression and tensile tests were included to accommodate the major loading behavior experienced by the tissue. The strains, strain rates, wait times, tare load, etc. were all obtained based on what the tissue experiences physiologically as well as what the testing equipment could reliably apply. The 2 hour recovery time between tests was deemed adequate as typically recovery time is usually similar to loading time37. The 1000 preconditioning cycles provided an analogous scenario to walking 1000 steps.

The displacement at 10g (0.1 N) force was used as reference during the test. Nonetheless, a 300 µm offset was adapted before all stages except finding reference length or thickness to capture the full range of loading and deformation included the unloaded state (Figure 5a). The 300 µm buffer allowed capturing of all the loading the sample goes through. The 10g load was used as an experimental reference location for application of desired strains based on pilot tests and load experienced by test samples during protocol optimization, the 10g (0.1 N) load allowed safe application of non destructive strains low enough to keep the tissue state in the beginning of the toe region.

2.6 Data acquisition
Force and displacement data were acquired at 2.5 kHz (maximum allowed by the testing machine, to capture high frequency components of data and the noise, which can be later removed as needed) with the exception of relaxation components in the stress relaxation tests. Our pilot tests demonstrated that the noise induced by the machine was prevalent during ramp loading but not when the machine was stationary, e.g. hold phase of stress-relaxation test. The hold data was acquired in four steps. With the first 10s acquired at full 2.5kHz, next 100s acquired at 1Hz, next 1000s acquired at 0.1Hz and the remaining data acquired at 0.01Hz. Video data for tensile tests was acquired in RAW format, 640 x 480 at 10 Hz only for stress-relaxation. Load data was filtered at 20Hz using a low pass third order Butterworth filter for the steps for establishing contact or obtaining reference length during the test. The preconditioning, pre- and post- preconditioning ramps and stress relaxation ramp load data was not filtered while testing so as to not miss the peaks.

2.7 Repeatability
Each of the entire test was repeated three times as it was deemed to be the lowest number of times a test can be repeated to be able to observe a trend, if any, while feasible to do so. For compression tests, each sample was tested for a total of six times over the course of three days. An unconfined compression test was followed by a rest period of two hours before the confined compression test. Tensile repeatability test set was conducted over two days with two tests on day one with two hours rest in between and one test the next day. Between two tests and after the day’s testing, the sample was kept in the fridge wrapped in a PBS soaked paper towel until the next test. For compression samples, thickness was measured before each test using OTMS. For tensile samples, thickness was measured once before the sample was placed in the grips. Grips were left attached to the tensile sample for the entire repeatability set and the sample was placed in the fridge after every test. Compression samples were completely removed from the test chambers and placed in the fridge after each test. A total of 54 tests were conducted for this study. Each sample was frozen after all the tests were completed so that it can be used if further testing is needed.

2.8 Data analysis and, quantification and assessment of repeatability
After analyzing the raw data, a reference state was decided by using a digital load threshold of 10g. This safely captures most of the toe region for all samples (may not maintain exact target strains, but that should have no effect on the linear region of the resultant stress-strain curve). The reason why the experimental and digital 10g displacement locations do not align is because the rate for finding the reference location was much lower than at which the cross-head reaches the same location when actual 20%/s strain rate is applied (Figure 5). All the data above the digital 10g location was cropped for analysis. The beginning of the curve was then adjusted to zero force – displacement. The force – displacement data were converted to engineering stress and strain using initial dimensions (Figure 5). During analysis, the force data for all ramps and preconditioning was filtered at 20 Hz to remove noise. It should be noted that the sign convention of the test system relies on compression being positive. For reporting purposes, all stresses and strains (compression or tension) are considered positive.

For full loading ramps before and after preconditioning (with target strains of 15%), high strain moduli were calculated using the top third of the stress-strain curve. Mean and standard deviations were calculated for each repeatability test set. For preconditioning, percent reductions in stress after 20 cycles and 1000 cycles were calculated to compare the extent of effect of number of cycles on stress reduction. For stress relaxation tests, average instantaneous modulus (AIM) for each test was calculated using the peaks at each target strain level of the test. Similarly, average equilibrium modulus (AEM) was calculated for each test using the equilibrium stress- strain values at each target strain level. AIM and AEM were calculated by fitting a straight line through the peak and equilibrium stress-strain values respectively. Mean and standard deviations were calculated for each repeatability test set. Coefficient of variation (COV) was calculated for each of these sets to assess variability within the set. Similarly, high strain instantaneous and equilibrium moduli were also calculated along with mean, standard deviations and coefficient of variation for each repeatability test set. High strain modulus was calculated by fitting a straight line to stress-strain data between 10-15% strain. Once the digital 10g (0.1 N) location was used, a re-assessment of deviation of achieved strains from target strains was also performed. Further, at each strain level the amount of relaxation (reduction in stress) was also calculated. It should be noted that these calculations were conducted to assess repeatability of testing protocols and provide gross macro-scale response. More elaborate constitutive modeling can be performed with the disseminated data (see Data Availability).

3. Results

Sample locations and reference dimensions are listed in Table 2. Each thickness measurement for compression tests is an average of thickness values obtained before each test, which in turn is average of ten measurements on each sample. Since thickness measurements for tension were taken only once for each set, they are averages of ten measurements on the sample. Reference lengths for tensile samples are averages of lengths obtained with 10g force before stress relaxation for each set and hence are averages of three measurements (one for each test in a set). All reference dimension measurements indicate minimal change in dimensions between repeatability tests (also see Appendix B).

From the ramps before and after preconditioning, high strain (upper third of the loading) moduli were calculated (Table 3). For preconditioning repeatability data, relaxation after 10 s (20 cycles) vs 1000 cycles was also calculated (Table 4). The stress relaxation after 10 s (20 cycles) for compression tests was within 11.36% to 27.16% and for tensile tests it was between 7.74% to 12.32% (Table 4). Similar behavior was observed for percentage relaxation at 1000 cycles with 46.89% to 79.40% for compression whereas, 20.59% to 29.94% for tensile samples. Figures 6 and 7 provide sample repeatability information for preconditioning peaks and pre- and post-preconditioning ramps, respectively, for a sample under unconfined compression taken from patella.

For the 18 repeatability sets (6 sets each for confined compression, unconfined compression, and uniaxial tension) with 3 tests in each set, average instantaneous (AIM) and equilibrium moduli (AEM) were calculated by fitting the peaks and equilibrium values respectively (Table 5) from the three step stress relaxation. Coefficient of variation (COV) was calculated for each of these sets to assess variability within the set (Table 5). Similarly, high strain average instantaneous and average equilibrium moduli were also calculated from the two highest strain levels in each test set (Table 6). Since, zero load-displacement application location was re-established using the digital 10g location, the potential variations in reference strains at each strain level were also calculated (Table 7). Average percent relaxation at each strain level after 30 min relaxation was also obtained (Table 8). Figure 8 provides a stress relaxation comparison for each test set for samples taken from patella and trochlear groove (Appendix C summarizes the rest of the stress relaxation tests). During tensile tests, image data was also obtained for stress relaxation which may be useful to calculate actual strain application and width changes (Figure 9). However, for this study, that analysis was not performed.

4. Discussion

Accurate and repeatable measurement of specimen-specific, regional, macro-scale mechanical properties (linear and nonlinear) for cartilage is critical to develop reliable FE models, which in turn, can guide implant and surgical procedures and aid understanding of the overall tissue mechanics. The specific goal of the study was to develop and provide a comprehensive protocol for characterizing articular cartilage, describing details that are not often presented in the relevant literature. Experimentation is a challenging endeavor and decisions of the experimenter and uncertainties associated with test protocols can lead to unexpected, unintended, or possibly erroneous results. This study addresses reproducibility of testing procedures, including assessing not just the protocol but the entire detailed workflow of mechanically testing a specific sample set. This information is critical for determining not just the usability of the developed procedures and protocol but also helps understanding the reliability of data for various uses of tissue characterization such as specimen specific finite element model development, the regional and biological variability in cartilage properties, or any decision making based on the material properties. This data set is also supported by specimen-specific geometry and joint level mechanical information (https://simtk.org/frs/?group_id=485, specimen: oks003) which can thus be used for specimen-specific computational model development. To the best of our knowledge, repeatability data for cartilage characterization has not been documented or published before and this presented protocol and assessment of procedures can provide a framework for future scientific studies.

The relatively low standard deviations in our reference dimension measurements (thickness for compression and length for tensile tests) (Table 2) indicate that reference dimension values, which were used for strain application, were consistent, thus minimizing the variation in strain application. The thicknesses of tensile samples were higher than typically tested in other studies and therefore a direct comparison may not be possible38, but low standard deviations in reference dimension measurements may also indicate that the samples recovered adequately between tests and structural damage was likely minimal.

For compression ramp loading before and after preconditioning, the moduli at higher strains (linear regions) after preconditioning were considerably lower than those before preconditioning. For tensile tests however, there was not as much of a difference. This indicates that recovery mechanisms in compression are likely different than for tension, which is an important consideration when developing and assessing protocols. The moduli for all tests for medial femoral condyle appear higher than those for other locations. Further, it appears that all samples relaxed more under compression (both confined and unconfined) compared to under tension after the application of 1000 cycles. From the stress relaxation tests, the standard deviations indicate that the intra-sample variability for relaxation was similar for the tests in each set. The equilibrium moduli appear to be within the range of previously reported studies12–14,28 (Table 5). It appears that the repeatability was not as good for confined compression tests as it was for either unconfined compression or uniaxial tension tests. This may be due to the issues with sample, filter and indenter alignment. As the 10g digital location was chosen as reference for zero load-displacement, deviations from the expected target strains were also calculated. The highest deviation from target strain appears to be for trochlear groove under unconfined compression.

The location dependent variability observed in the material properties (albeit just one specimen in this report) indicates that use of literature for either characterization protocols or material properties should be carefully considered when using in finite element analysis or other computational modeling strategies. This study also demonstrates some of the challenges associated with tissue material characterization with results underscoring the need for specimen-specific material properties where testing protocols are described and thoroughly evaluated in concert with computational studies that show similar effects of material properties on finite element model predictions8,16. Pragmatically it is not always possible to obtain specimen-specific properties for a given specimen. For material properties, characterization data with vetted procedures can provide improved reliability in modeling/simulation outputs. The wide range of location dependent moduli values reported here also show agreement with results from previous studies28. This emphasizes the potential need to provide and assess multi-region, location specific material properties. Along with utility in computational modeling, characterization of cartilage mechanical behavior reliably done can be useful in understanding how variable the behavior is, within a given specimen depending on the location, within a population, for different species, as well as in healthy and diseased states.

The tests were repeated in their entirety for compression tests and for tensile tests (except thickness measurement). This encompasses both the repeatability of the test as well as reproducibility of the entire experimental procedure. This ensured capturing the effects due to experimental procedures, protocols, potential changes in tissue geometry or degradation, if any. The quality of the cartilage tissue was also assessed by a trained radiologist and deemed healthy, providing confidence that the tests were not conducted on a readily degraded tissue. It is important to note that the tissue samples recovered to original dimensions at each test and showed no visible signs of failure or deterioration, an indication that they were not damaged as part of the application of desired strain levels and/or strain rates. Additionally, tissue samples also recovered to original force-displacement behavior as evident from the repeatable first ramp loading at the beginning of each test. Samples also demonstrated repeatable viscoelastic response, e.g. during stress relaxation after completion of preconditioning.

Some limitations are noted. First, this study was conducted on cartilage samples taken from one specimen. However, our focus in this report was maintaining and reporting comprehensive protocol control and on the acquired repeatability data. Samples from major locations of the knee were taken to be as comprehensive in our inclusion of sample locations as possible given the available resources. Also, due to feasibility and limited size of cartilage available, only one compression and one tensile sample was taken from each major location. This may limit the comparison between complementary locations, e.g., patella and trochlear groove, since samples may not fall on exact opposite load bearing locations. Tensile samples were also thicker and may have included not just the superficial zone. However, the data for tensile samples at least indicates the adequacy of the testing protocols to provide reasonably repeatable results. The donor was relatively young and hence, the repeatability may be better due to the perceived health of the tissue. Comprehensive testing of older or diseased / osteoarthritis specimens may help establish experimentation uncertainties and if needed, tailor the protocols for a wide range of age and disease states. The standardization of the protocol was done on limited samples and thus the protocol may still have room for improvement. Another potential concern was the potential effect of multiple freeze-thaw cycles on the tissue behavior. To address this, a test sample was tested under unconfined compression for up to seven freeze-thaw cycles before any visible deterioration was observed. This provided confidence that the two freeze thaw cycles that the samples were subjected to (specimen procured frozen from vendor and dissected tissue frozen until testing day) did not significantly affect the tissue behavior. In the future, the authors would like to extend this study to multiple locations from the same specimen and multiple specimens with varying condition and tissue types. Further assessment of micro structures of specimens via MRI or histology may be done to evaluate the effects of fiber orientation, etc., on both the repeatability and material properties.

The initial driving motivation for developing this characterization protocol and repeatability assessment was to be able to obtain reliable material properties for specimen-specific FE models of the knee. This study focused entirely on development of these protocols for articular cartilage and assessing their quality. The data, which are publicly disseminated, now can be used to not only fit three dimensional (3D) material representations but also to streamline and develop other tissue characterization protocols and addressing questions such as regional variations in material properties. For this study, mechanical metrics representative of gross tissue behavior were used to evaluate repeatability. Therefore, the reported material properties are rudimentary with the sole purpose of assessing repeatability. Now that the repeatability of the experiment, therefore the uncertainty of data, has been quantified, the future focus will be to obtain comprehensive representations of tissue response, by the authors of this study or any other interested parties. Fidelity of constitutive models, e.g., compressibility, viscoelasticity, can be evaluated using multiple available sets for different loading modes. Tension compression nonlinearity39, biphasic response40, among many other mechanical nonlinearities of cartilage, can potentially be represented using this comprehensive dataset, which includes detailed tests of samples from the same region and multiple cartilage regions from the same knee. These material models can be incorporated in 3D FE models where complicated mechanical loading environments may be simulated. Additionally, conditions such as cartilage behavior before and after preconditioning can also be simulated. The data sets provide the opportunity to fit constitutive models to the whole temporal loading pattern of testing protocol or its portions. One may use only stress relaxation data sets to fit biphasic or viscoelastic models; another may utilize only loading response to fit nonlinear elastic models; yet another may decide to explore the role of pre-conditioning by fitting models representative of such transient behavior. Many different types of models can be evaluated to capture both long term and short term time-dependent and nonlinear response of the tissue. Consequently, this data set may serve as an opportunity to establish benchmark cases to test cartilage constitutive models in terms of their fidelity to represent both spatial and temporal loading modes. Further, when a constitutive model is built that represents the entire tissue behavior and the fit error is as good as the error in repeatability, this experiment can be replicated in silico.

Detailed knowledge of the protocols and the testing errors associated with this set establish a comprehensive reference for future testing and provide a foundation for future mechanical characterization studies in development. Knowledge of these detailed procedures can be highly valuable to future researchers in cartilage (and whole knee) biomechanical testing and analyses. These protocols can be optimized further to improve the reproducibility as well as develop pipelines to streamline tissue characterization for other tissues. Even though the driving motivation for this study was the need of specimen-specific and location dependent material information for finite element analysis of the knee, the impact of these data extends beyond this goal. Standardization of experimental procedures is useful for producing more comparable and consistent results across studies. The results presented here underscore the need to perform, thoroughly evaluate, and report the tissue characterization protocols to better understand both results of experiments and use of measured parameters in advanced computational models.

Figures

See figures document

Tables

See tables document
Acknowledgments

This study was part of the Open Knee(s) project, which was funded by the National Institute of General Medical Sciences, National Institutes of Health (NIH) under Grant R01GM104139 (Principal Investigator: Ahmet Erdemir). Insights provided by the Open Knee(s) Advisory Board (Jack Andrish, Morgan Jones, Paul Saluan, Carl Winalski, Yasin Dhaher, Trent Guess and Rami Korhonen) are greatly appreciated. Authors would also like to thank Craig Bennetts for helping with specimen dissection. Up-to-date information on Open Knee(s) project can be found at https://simtk.org/projects/openknee. Part of the support for continuing the work also came from Grant No. R01EB024573 funded by the National Institute of Biomedical Imaging and Bioengineering, National Institutes of Health.

Competing interests

Snehal Chokhandre has no competing interests. Ahmet Erdemir owns and operates innodof, LLC, a consulting company for modeling and simulation.

Author contributions

Conceived and designed the experiments: Snehal Chokhandre and Ahmet Erdemir. Performed experiments: Snehal Chokhandre. Analyzed the data: Snehal Chokhandre. Wrote the paper: Snehal Chokhandre, Ahmet Erdemir.

References

1.	Goldsmith, A. A., Hayes, A. & Clift, S. E. Application of finite elements to the stress analysis of articular cartilage. Med Eng Phys 18, 89–98 (1996).
2.	Guo, H., Maher, S. A. & Torzilli, P. A. A biphasic finite element study on the role of the articular cartilage superficial zone in confined compression. J Biomech 48, 166–170 (2015).
3.	Clift, S. E. Finite-element analysis in cartilage biomechanics. J Biomed Eng 14, 217–221 (1992).
4.	Halonen, K. S. et al. Deformation of articular cartilage during static loading of a knee joint--experimental and finite element analysis. J Biomech 47, 2467–2474 (2014).
5.	Mononen, M. E., Liukkonen, M. K. & Korhonen, R. K. Utilizing Atlas-Based Modeling to Predict Knee Joint Cartilage Degeneration: Data from the Osteoarthritis Initiative. Ann Biomed Eng 47, 813–825 (2019).
6.	Kazemi, M., Dabiri, Y. & Li, L. P. Recent Advances in Computational Mechanics of the Human Knee Joint. Computational and Mathematical Methods in Medicine (2013). doi:10.1155/2013/718423
7.	Kiapour, A. et al. Finite element model of the knee for investigation of injury mechanisms: development and validation. J Biomech Eng 136, 011002 (2014).
8.	Peters, A. E., Akhtar, R., Comerford, E. J. & Bates, K. T. Tissue material properties and computational modelling of the human tibiofemoral joint: a critical review. PeerJ 6, e4298 (2018).
9.	Blankevoort, L., Kuiper, J. H., Huiskes, R. & Grootenboer, H. J. Articular contact in a three-dimensional model of the knee. J Biomech 24, 1019–1031 (1991).
10.	Abo-Alhol, T. R. et al. Patellar mechanics during simulated kneeling in the natural and implanted knee. Journal of Biomechanics 47, 1045–1051 (2014).
11.	Nissi, M. J. et al. Estimation of mechanical properties of articular cartilage with MRI - dGEMRIC, T2 and T1 imaging in different species with variable stages of maturation. Osteoarthr. Cartil. 15, 1141–1148 (2007).
12.	Boschetti, F., Pennati, G., Gervaso, F., Peretti, G. M. & Dubini, G. Biomechanical properties of human articular cartilage under compressive loads. Biorheology 41, 159–166 (2004).
13.	Jurvelin, J. S., Buschmann, M. D. & Hunziker, E. B. Mechanical anisotropy of the human knee articular cartilage in compression. Proc Inst Mech Eng H 217, 215–219 (2003).
14.	Treppo, S. et al. Comparison of biomechanical and biochemical properties of cartilage from human knee and ankle pairs. J. Orthop. Res. 18, 739–748 (2000).
15.	Armstrong, C. G. & Mow, V. C. Variations in the intrinsic mechanical properties of human articular cartilage with age, degeneration, and water content. J Bone Joint Surg Am 64, 88–94 (1982).
16.	Yao, J., Funkenbusch, P. D., Snibbe, J., Maloney, M. & Lerner, A. L. Sensitivities of medial meniscal motion and deformation to material properties of articular cartilage, meniscus and meniscal attachments using design of experiments methods. J Biomech Eng 128, 399–408 (2006).
17.	Bursać, P. M., Obitz, T. W., Eisenberg, S. R. & Stamenović, D. Confined and unconfined stress relaxation of cartilage: appropriateness of a transversely isotropic analysis. Journal of Biomechanics 32, 1125–1130 (1999).
18.	Lu, X. L. & Mow, V. C. Biomechanics of articular cartilage and determination of material properties. Med Sci Sports Exerc 40, 193–199 (2008).
19.	Chen, X., Zimmerman, B. K. & Lu, X. L. Determine the equilibrium mechanical properties of articular cartilage from the short-term indentation response. J Biomech 48, 176–180 (2015).
20.	Morel, V., Berutto, C. & Quinn, T. M. Effects of damage in the articular surface on the cartilage response to injurious compression in vitro. J Biomech 39, 924–930 (2006).
21.	Töyräs, J. et al. Estimation of the Young’s modulus of articular cartilage using an arthroscopic indentation instrument and ultrasonic measurement of tissue thickness. J Biomech 34, 251–256 (2001).
22.	Langelier, E. & Buschmann, M. D. Increasing strain and strain rate strengthen transient stiffness but weaken the response to subsequent compression for articular cartilage in unconfined compression. J Biomech 36, 853–859 (2003).
23.	Thambyah, A., Nather, A. & Goh, J. Mechanical properties of articular cartilage covered by the meniscus. Osteoarthr. Cartil. 14, 580–588 (2006).
24.	Athanasiou, K. A., Rosenwasser, M. P., Buckwalter, J. A., Malinin, T. I. & Mow, V. C. Interspecies comparisons of in situ intrinsic mechanical properties of distal femoral cartilage. J. Orthop. Res. 9, 330–340 (1991).
25.	Li, L. P., Buschmann, M. D. & Shirazi-Adl, A. Strain-rate dependent stiffness of articular cartilage in unconfined compression. J Biomech Eng 125, 161–168 (2003).
26.	Danso, E. K., Honkanen, J. T. J., Saarakkala, S. & Korhonen, R. K. Comparison of nonlinear mechanical properties of bovine articular cartilage and meniscus. J Biomech 47, 200–206 (2014).
27.	Korhonen, R. K. & Jurvelin, J. S. Compressive and tensile properties of articular cartilage in axial loading are modulated differently by osmotic environment. Med Eng Phys 32, 155–160 (2010).
28.	Deneweth, J. M., Newman, K. E., Sylvia, S. M., McLean, S. G. & Arruda, E. M. Heterogeneity of tibial plateau cartilage in response to a physiological compressive strain rate. J. Orthop. Res. 31, 370–375 (2013).
29.	Charlebois, M., McKee, M. D. & Buschmann, M. D. Nonlinear tensile properties of bovine articular cartilage and their variation with age and depth. J Biomech Eng 126, 129–137 (2004).
30.	Gao, L.-L., Zhang, C.-Q., Gao, H., Liu, Z.-D. & Xiao, P.-P. Depth and rate dependent mechanical behaviors for articular cartilage: experiments and theoretical predictions. Mater Sci Eng C Mater Biol Appl 38, 244–251 (2014).
31.	Lai, J. H. & Levenston, M. E. Meniscus and cartilage exhibit distinct intra-tissue strain distributions under unconfined compression. Osteoarthr. Cartil. 18, 1291–1299 (2010).
32.	Ahsanizadeh, S. & Li, L. Visco-hyperelastic constitutive modeling of soft tissues based on short and long-term internal variables. Biomedical engineering online 14, 29 (2015).
33.	Cheng, S., Clarke, E. C. & Bilston, L. E. The effects of preconditioning strain on measured tissue properties. J Biomech 42, 1360–1362 (2009).
34.	Ebrahimi, M., Mohammadi, A., Ristaniemi, A., Stenroth, L. & Korhonen, R. K. The effect of different preconditioning protocols on repeatability of bovine ACL stress-relaxation response in tension. J Mech Behav Biomed Mater 90, 493–501 (2019).
35.	Hosseini, S. M., Wilson, W., Ito, K. & van Donkelaar, C. C. How preconditioning affects the measurement of poro-viscoelastic mechanical properties in biological tissues. Biomech Model Mechanobiol 13, 503–513 (2014).
36.	Gu, K. B. & Li, L. P. A human knee joint model considering fluid pressure and fiber orientation in cartilages and menisci. Medical Engineering & Physics 33, 497–503 (2011).
37.	Shepherd, D. E. & Seedhom, B. B. A technique for measuring the compressive modulus of articular cartilage under physiological loading rates with preliminary results. Proc Inst Mech Eng H 211, 155–165 (1997).
38. Akizuki, S. et al. Tensile properties of human knee joint cartilage: I. Influence of ionic conditions, weight bearing, and fibrillation on the tensile modulus. J. Orthop. Res. 4, 379–392 (1986).
39.	Soltz, M. A. & Ateshian, G. A. A Conewise Linear Elasticity mixture model for the analysis of tension-compression nonlinearity in articular cartilage. J Biomech Eng 122, 576–586 (2000).
40.	Huang, C. Y., Mow, V. C. & Ateshian, G. A. The role of flow-independent viscoelasticity in the biphasic tensile and compressive responses of articular cartilage. J Biomech Eng 123, 410–417 (2001).

Data availability

A static version of raw data (as utilized for this publication) is available at https://doi.org/10.18735/WTJZ-N328 . All the protocols, testing iterations, raw data, analysis scripts, and an evolving set of analysis outcomes are also available at the Open Knee(s) project site, https://simtk.org/projects/openknee.